## Slide 1: Using the 60 Channel MEA
Build 3.6.5 (04/22/2012)

## Slide 2: Disclaimer
This data acquisition system is a work in progress.
The executable file is limited in scope and will be continuously upgraded
Any problems or questions, call Mary at 603-674-8882 or email to mtgjbird@hotmail.com

## Slide 3: The Steps to Collecting Data
Sample Preparation
DAQ system set up
Recording Chamber set up
Microscope set up
Collecting Data (see last slide for abbreviated steps)
Exporting Data
Shutting down

## Slide 4: Step 1: Sample Preparation
Clean the glass MEA recording chamber (according to manufacturer’s instructions) and disinfect (with 70 - 95% ethanol).
Coat the interior surface as needed with PLL or substrate of choice (indicated by the cell lines or tissue slice).
Perform all prep under aseptic conditions (ie in the hood, using sterile instruments and media).
Cover the recording chamber with sterile filter while transporting to the microscope and keep covered until recording.

## Slide 5: Step 2: DAQ Instrument Set Up
Check the connections to the various components:
  Plug the three wires from the PS20W (blue, red, and black) into the corresponding ports on the SCB-100 connector box
  Plug the 68-pin plug from the SCB-100 into the receptacle on the MEA recording amplifier/signal conditioner
  Plug the USB cord from the blue eyepiece camera on the microscope into a USB port on the Dell Studio XPS computer. (This wire should come out of the camera in the direction of the door to the lab so that the sample images are aligned properly [this will be corrected programmatically at a later time]).

## Slide 6: Step 2: DAQ Instrument Set Up (Stimulation System)
Plug the USB from the DT9814 (A) to a USB port on the computer.
There are two wires (one red, one black) attached to the “DAC0” and “DAC0 return” on the left side. They will connect to the MEA as follows: Connect the black spring hook (B) to the channel you want to stimulate on the MEA (C); Clip the alligator clip to another spring hook and clip that one to the channel next to the channel you are stimulating (D). For example, if you want to stimulate the electrode at 3,8, you would clip the second spring hook to 3,7 or 2,8, or 4,8
A
B
C
D

## Slide 7: Step 3: Setting up Recording Chamber
Place the glass MEA recording chamber into the MEA amplifier/signal conditioner so that the thick line indicated in picture points away from you (towards the “G” that is written on the lever described in the next section).
Look at the underside of the top of the unit and make sure all pins are fully extended (they are spring-loaded and sometimes get stuck – if stuck, pull gently into position.) Carefully lower the top of the unit onto the glass MEA and lock using the 2 small levers.
Use care to avoid splashing media or buffer on the outside of the recording area on the glass MEA as this corrodes the electrodes.

## Slide 8: Step 4: Setting up the Microscope (if acquiring images)
Place the MEA amplifier onto the microscope stage taking care to not break any wires. The 68 pin plug should be to the left side of the amplifier.
Turn on the microscope and move the stage so that the MEA is visible through the eyepieces. The knob on the right side of the Microscope should be set to “BI” or “OM.” Adjust the focus as necessary.

## Slide 9: Step 5: Collecting Data
Turn on the outlet strip and the PS20W (the switch is on the backside of the unit).
Open the “Raptor 3.6.5” program by dbl-clicking on the Hawk icon on the desk top

## Slide 10: Step 5: Collecting Data (continued)
You should see a screen that looks like this:

## Slide 11: Step 5: Collecting Data (control tabs)
When you open Raptor 3.6.5, there are 4 tabs available to set up your study. The first tab, “60 Channels” is where you set up the acquisition time and collect all channels available, the 2nd tab, “Stimulation” is for electrically stimulating a channel on the MEA with a user-defined voltage and waveform and the DT9814, the 3rd tab is for collecting Images (camera must be connected to USB port for this ), and the 4th tab is a large text box where you can add sample info. The 2nd, and 3rd tabs each have a button associated with them that must be clicked in order to use those functions.

## Slide 12: Step 5: Collecting Data (no stimulation)
Choose an acquisition time (type it in or use the arrows).
Click “Operate” and choose “Run”

## Slide 13
Step 5: Collecting Data
(with Stimulation)
In the “60 Channels” tab, Choose an acquisition time (type it in or use the arrows) then, click the button next to “Stimulation”

## Slide 14: Step 5: Collecting Data (with Stimulation)
Click on the “Stimulation” Tab to get the following screen. Choose the time delay for the stimulation (this is the time after the start of data acquisition), then choose the length of time you want to stimulate (in seconds). If you want the Glutamate biological signal, it is set by default but if you want to choose a different waveform, click the toggle switch to point to the right and choose sine, square, triangle, or sawtooth, then set the maximum voltage (it is a biphasic waveform so if you enter 0.1V, the range is -0.1 to 0.1V), and select the time – note that the current settings have one cycle = 0.66 seconds. Maximum output is 2.4V so you will see artifacts if voltage greater than that is chosen.

## Slide 15: Step 5: Collecting Data (with Stimulation)
Go Back to the “60 Channels” tab, Click “Operate” and choose “Run”

## Slide 16: Step 5: Collecting Data (with Stimulation)
If you forget to plug in the DT9814, you will get an error that looks something like this:

## Slide 17: Step 5: Collecting Data (continued)
If you need to stop it before the chosen acquisition time is up, press the stop sign at any time (data will not save once the stop sign has been pressed).

## Slide 18: Step 6: Saving the Data
After the data has been collected, it will flash once and the background will automatically subtract. Click on the dark green “click to save” button (it will flash bright green) .

## Slide 19: Step 6: Saving the Data (continued)
Choose a unique filename that contains enough information so you can recognize the file and type of study. Enter it into the box and click “OK.” This will happen twice – one file is an LVM file and the other is a TDMS file (for streaming).

## Slide 20: Step 6: Saving the Data (continued)
Click the stop sign after data has been saved

## Slide 21: Step 7: Collecting Data from other samples
Click “Edit” and choose “Reinitialize Values to Default” then repeat steps 5 and 6 for the next samples

## Slide 22: Step 8: Looking at Your Data
To get a closer look at any graph, right-click on it and choose “visible Items” and then “Graph Palette.” This give you the tools to zoom or stretch a graph while still in the Raptor program
Zoom control
Zoom opens this

## Slide 23: Other Features
Click the “Sample Info” Tab and you can insert text that describes the experiment or the samples. The output is currently a picture of the text. Later versions will have it exported as text.

## Slide 24: Step 9: Getting your Data (continued)
Individual graphs can be exported individually into Excel or as Images by right-clicking on the graph and choosing “Export”

## Slide 25: Step 9: Getting your Data (continued)
To get your image, click the Image Tab and then right click on the image, choose “Copy Data” and paste it (control+V) into Paint and save it. It then can be opened with any image analysis software (Photoshop Elements is shown).

## Slide 26: Step 10: Getting your Data (in EXCEL)
The data is saved automatically as an “.lvm” file – this is a kind of text file and the complete data set (excluding the images) can be opened into Excel using the import from text file option.
Open Excel, Click “Data” and choose “from text.” You will need to find the correct data file and change file type to “all files”
The file will be formatted with one X-column (time) and 60 Y-columns (each channel, labeled 0-59)

## Slide 27: Getting your Data (in DIAdem)
DIAdem is a National Instruments program that is specifically designed to handle LabView files.
See “Raptor Analysis” powerpoint for details on how to use it

## Slide 28: Step 11: Shutting Down
Check to make sure your data files collected and saved correctly.
Shut off the PS20W.
Remove the glass MEA recording chamber from the amplifier.
Shut off the microscope.
Shut down the computer.
Shut off the outlet strip

## Slide 29: Abbreviated Data Collection Steps
Turn on PS20W
Click Hawk Icon (“Raptor 3.6.5”)
Choose an acquisition time
If stimulating, click button on stimulation tab, click stimulation tab to open that page, enter values for time delay, waveform type, and voltage.
Click “60 channels” tab
Click “Operate” and choose “Run”
Click the “click to save” button and choose a unique filename, save twice.
Press the stop button to stop the program after the data is saved (or any time you need to end an acquisition early)

## Slide 30: Extra Slides: Program and Hardware set-up.

## Slide 31: Hardware Set-up
DAQ card is National Instruments PCI-6033E
Driver is NI-DAQmx

## Slide 32: NI PCI-6033E Accuracy

## Slide 33: Wiring Diagram
PCI-6033E to SCB-100 connected via SH100100 cable
Wiring within SCB-100

## Slide 34: Multichannel Systems Amplifier and Signal Conditioner
PS20W Power supply powers the MEA 1060-Inv amplifier
Wiring Diagram

## Slide 35: Raptor 3.6.5 Program Set-up
Raptor was created using LabView 2011 SP1
Raptor 3.6.5 was built from LV project 022012.lvproj, testrecording022012, Hawkandlightnin2 Icon 48x48.
To rebuild an executable, the following modifications are necessary:
  under “source file settings”, click testrecording022012 then click “customize VI properties,” then uncheck “run when opened,”
  Under “advanced”, uncheck “copy error codes”

## Slide 36: Block Diagram of Raptor 3.6.5
Acquisition Time
DAQ Assistant
Background subtraction
Bandpass filter 0.5 – 200Hz
Camera Control (IMAQ/Vision Express)
Front Page Graphs
Data Written to .lvm file (tab-delimited text file) – there are now 2 of these – the second writes to TDMS files
Tab Controls (for front page)
Stimulus program on next 2 slides

## Slide 37: Stimulus Program Biological

## Slide 38: Stimulus Program Sine/Square/Triangle/Sawtooth

## Slide 39: DAQ Assistant Set-up
Software triggering, internal clock

## Slide 40: Camera Set-up (IMAQ/Vision Express)
Cam2: TCAM3 has Direct Show Driver
Raptor 3.6.5 will collect a single snapshot at start of data collection. Continuous collection can be added later, if needed.
Additional settings within Raptor 3.6.5

## Slide 41: Estimated DAQ Hardware Cost (savings of $2798.03)
Total: $538.97
(includes camera)
Total: $3337.00 (without camera)
New
Used