## Slide 1: Using the 60 Channel MEA
Build 0703 (2016)

## Slide 2: Disclaimer
This data acquisition system is a work in progress.
The executable file is limited in scope and will be continuously upgraded
Any problems or questions, email to mtgjbird@hotmail.com

## Slide 3: The Steps to Collecting Data
Sample Preparation
DAQ system set up
Recording Chamber set up
Microscope set up
Collecting Data (see last slide for abbreviated steps)
Exporting Data
Shutting down

## Slide 4: Step 1: Sample Preparation
Clean the glass MEA recording chamber (according to manufacturer’s instructions) and disinfect (with 70 - 95% ethanol).
Coat the interior surface as needed with PLL or substrate of choice (indicated by the cell lines or tissue slice).
Perform all prep under aseptic conditions (ie in the hood, using sterile instruments and media).
Cover the recording chamber with sterile filter while transporting to the microscope and keep covered until recording.

## Slide 5: Step 2: DAQ Instrument Set Up
Check the connections to the various components:
  Plug the three wires from the PS20W (blue, red, and black) into the corresponding ports on the SCB-100 connector box
  Plug the 68-pin plug from the SCB-100 into the receptacle on the MEA recording amplifier/signal conditioner
  Plug the USB cord from the blue eyepiece camera on the microscope into a USB port on the Dell Studio XPS computer. (This wire should come out of the camera in the direction of the door to the lab so that the sample images are aligned properly [this will be corrected programmatically at a later time]).

## Slide 6: Step 3: Setting up Recording Chamber
Place the glass MEA recording chamber into the MEA amplifier/signal conditioner so that the thick line indicated in picture points away from you (towards the “G” that is written on the lever described in the next section).
Look at the underside of the top of the unit and make sure all pins are fully extended (they are spring-loaded and sometimes get stuck – if stuck, pull gently into position.) Carefully lower the top of the unit onto the glass MEA and lock using the 2 small levers.
Use care to avoid splashing media or buffer on the outside of the recording area on the glass MEA as this corrodes the electrodes.

## Slide 7: Step 4: Setting up the Microscope (if acquiring images)
Place the MEA amplifier onto the microscope stage taking care to not break any wires. The 68 pin plug should be to the left side of the amplifier.
Turn on the microscope and move the stage so that the MEA is visible through the eyepieces. The knob on the right side of the Microscope should be set to “BI” or “OM.” Adjust the focus as necessary.

## Slide 8: Step 5: Collecting Data
Turn on the outlet strip and the PS20W (the switch is on the backside of the unit).
Open the “Raptor” program by dbl-clicking on the Hawk icon on the desk top

## Slide 9: Step 5: Collecting Data (continued)
You should see a screen that looks like this:
You may need to click “Window” then “full size” to make it fit your screen

## Slide 10: Step 5: Collecting Data (control tabs)
When you open Raptor, there are tabs available to set up your study. The first tab, “60 Channels” is where you set up the acquisition time and collect all channels available, the 2nd tab, is for collecting Images (camera must be connected to USB port for this ), and the 3rd tab is a large text box where you can add sample info.

## Slide 11: Step 5: Collecting Data
Choose an acquisition time (type it in or use the arrows).
Click “Operate” and choose “Run” or click on the right arrow

## Slide 12: Step 5: Collecting Data (continued)
If you need to stop it before the chosen acquisition time is up, press the stop sign at any time (data will not save once the stop sign has been pressed).

## Slide 13: Step 6: Saving the Data
After the data has been collected, it will flash once and the background will automatically subtract. Click on the dark green “click to save” button (it will flash bright green) .

## Slide 14: Step 6: Saving the Data (continued)
Choose a unique filename that contains enough information so you can recognize the file and type of study. Enter it into the box and click “OK.” This will happen twice – one file is an LVM file and the other is a TDMS file (for streaming).

## Slide 15: Step 6: Saving the Data (continued)
Click the stop sign after data has been saved

## Slide 16: Step 7: Collecting Data from other samples
Click “Edit” and choose “Reinitialize Values to Default” then repeat steps 5 and 6 for the next samples

## Slide 17: Step 8: Looking at Your Data
To get a closer look at any graph, right-click on it and choose “visible Items” and then “Graph Palette.” This give you the tools to zoom or stretch a graph while still in the Raptor program
Zoom control
Zoom opens this

## Slide 18: Other Features
Click the “Sample Info” Tab and you can insert text that describes the experiment or the samples. The output is currently a picture of the text. Later versions will have it exported as text.

## Slide 19: Step 9: Getting your Data (continued)
Individual graphs can be exported individually into Excel or as Images by right-clicking on the graph and choosing “Export”

## Slide 20: Step 9: Getting your Data (continued)
To get your image, click the Image Tab and then right click on the image, choose “Copy Data” and paste it (control+V) into Paint and save it. It then can be opened with any image analysis software (Photoshop Elements is shown).

## Slide 21: Step 10: Getting your Data (in EXCEL)
The data is saved automatically as an “.lvm” file – this is a kind of text file and the complete data set (excluding the images) can be opened into Excel using the import from text file option.
Open Excel, Click “Data” and choose “from text.” You will need to find the correct data file and change file type to “all files”
The file will be formatted with one X-column (time) and 60 Y-columns (each channel, labeled 0-59)

## Slide 22: Step 11: Shutting Down
Check to make sure your data files collected and saved correctly.
Shut off the PS20W.
Remove the glass MEA recording chamber from the amplifier.
Shut off the microscope.
Shut down the computer.
Shut off the outlet strip

## Slide 23: Abbreviated Data Collection Steps
Turn on PS20W
Click Hawk Icon
Choose an acquisition time
Click “Operate” and choose “Run”
Click the “click to save” button and choose a unique filename, save twice.
Press the stop button to stop the program after the data is saved (or any time you need to end an acquisition early)

## Slide 24: Extra Slides: Program and Hardware set-up.

## Slide 25: Hardware Set-up
DAQ card is National Instruments PCI-6033E
Driver is NI-DAQmx

## Slide 26: NI PCI-6033E Accuracy

## Slide 27: Wiring Diagram
PCI-6033E to SCB-100 connected via SH100100 cable
Wiring within SCB-100

## Slide 28: Multichannel Systems Amplifier and Signal Conditioner
PS20W Power supply powers the MEA 1060-Inv amplifier
Wiring Diagram

## Slide 29: Raptor Program Set-up
Raptor was created using LabView 2013 SP1
Raptor was built from LV project 07032016.lvproj, testrecording07032016, Hawklightningdish Icon 48x48.
To rebuild an executable, the following modifications are necessary:
  under “source file settings”, click testrecording07032016 then click “customize VI properties,” then uncheck “run when opened,”
  Under “advanced”, uncheck “copy error codes”

## Slide 30: Block Diagram of Raptor
Acquisition Time
DAQ Assistant
Background subtraction
Bandpass filter 0.5 – 200Hz
Camera Control (IMAQ/Vision Express)
Front Page Graphs
Data Written to .lvm file (tab-delimited text file) – there are now 2 of these – the second writes to TDMS files
Tab Controls (for front page)
Stimulus program on next 2 slides

## Slide 31: DAQ Assistant Set-up
Software triggering, internal clock

## Slide 32: Camera Set-up (IMAQ/Vision Express)
Cam2: TCAM3 has Direct Show Driver
Raptor will collect a single snapshot at start of data collection. Continuous collection can be added later, if needed.
Additional settings within Raptor

## Slide 33: Estimated DAQ Hardware Cost (savings of $2798.03)
Total: $538.97
(includes camera)
Total: $3337.00 (without camera)
New
Used